CAD 3D printable device file: 3D Printed Embouchure Assist Device for Saxophone Players with Facial Palsy

Media 1: Video of patient playing the tenor saxophone without the device

Media 2: Video of patient playing the tenor saxophone with the device

Table S1. Device iteration phases and descriptions.

Phase 1: Initial Prototypes: U-cup and wall designs

Iteration Phase 1 Designs: Four initial designs were created before the first meeting with patient. Design 1: Double wall seal Design 2: Tri-directional wall Design 3: Symmetric bilateral U-Cup seal Design 4: Asymmetric unilateral U-Cup seal Images of patient with iteration phase 1 designs. Images 1 and 2 show designs 2 and 3, respectively. Photographs were taken to visualize the affected side of the patient and the degree to which an air seal was effectively able to be created with the embouchure. Description: These first prototype of the device, the tri-directional wall, was designed based on the 3D scans of the patients face and oral measurements to block air leakage from the patient’s mouth. However, during prototype trials with the patient, the wall proved to be too distally positioned on the mouthpiece to provide enough counter-pressure to stabilize the lip and additionally did not extend laterally enough to seal the gap where air was leaking. In analyzing patient as he played without the device, it was noted that his upper lip was unable to tighten around the mouthpiece to create a seal. Thus, the next prototypes, integrating bilateral and unilateral U-cup seals, were created with the intent of holding the patient's lip in place as he played. Finite Element Method (FEM) analysis showed that a device which held the patient’s lip in a groove, like that in the U-cup designs, would allow the bottom of the upper lip to withstand more pressure without breaking its seal with the mouthpiece (lower picture), as compared to the lip laying on the flat surface of the mouthpiece alone (upper picture). On trial in person, it quickly proved itself to be unusable, as it widened the mouthpiece volume around which the lip had to form a seal, making an effective embouchure even more difficult to achieve. FEM Analysis of Lip on Mouthpiece, with lip on straight flat surface (above) and lip fitted into a U-cup groove (below).

Phase 2: Adjusted U-cup seals and tri-directional walls

Iteration Phase 2 Designs: Updated designs created after the initial meeting with the patient. Design 1: asymmetric U-Cup seal but thinner towards mouthpiece Design 2: symmetric U-Cup seal but thinner towards mouthpiece Design 3: full frontal cover moved towards tip of mouthpiece. Description: Phase 2 of designs aimed to reduce the volume of the U-Cup seal, as well as to bring the wall support more proximally along the mouthpiece so that it would press against the patient’s face with more counter-pressure. Results were similar to those in phase 1, and an effective seal was unable to be created.

Phase 3: Extra-oral Flap

Iteration Phase 3 Designs: Iteration Phase 3 Designs: Design 1: asymmetric solid wall Design 2: asymmetric curved wall with increased thickness of 0.2in Design 3: asymmetric solid wall with increased curvature angle Design 4: asymmetric solid wall with increased curvature angle and increased thickness of 0.2in Images of patient playing with phase 3 of the prototpye, Images 1 and 2 show designs 1 and 3, respectively. Image 3 shows design 2 Description: This next prototype, a unilateral extraoral flap, mimicked a metal unilateral attachment created in a previous study18 except out of 3D printed material. The patient did not feel as though this device applied enough pressure to help him make a seal with the mouthpiece.

Phase 4: Intra-oral support piece with extra-oral flap

Iteration Phase 4 Designs: Design 1: small intraoral piece with interface surface on a 30-degree plane towards the face and 20-degree curved wall. Design 2: small intraoral piece with interface surface on a 35-degree plane and 20-degree curved wall. Design 3: small intraoral piece with interface surface on a 40-degree plane and 20-degree curved wall. Design 4: small intraoral piece with interface surface on a 40-degree plane and 30-degree curved wall. Pictures of the patient attempting to produce sound with this iteration of the device in place, with an intra-oral and extra-oral support system. Description: In this next round of iterations, a small intraoral component was also incorporated for additional support, intended to fill in the gap in the patient's lips where air was escaping due to his facial palsy. The patient found that the intraoral component hindered his ability to close his mouth and made it more difficult to create a proper seal around the mouthpiece. The area of the flap was increased from past iterations to better cover the area of air leakage, and optimal angle was experimented with.

Phase 5: Intra-oral support piece with extra-oral flap and thicker support rod (trial 2)

Iteration Phase 5 Designs: Design 1: 0.75x intraoral area from phase 4, 0.5x vertical area wall from phase 4, 30-degree curve Design 2: 0.75x intraoral, 2x vertical area wall, 20-degree curve. Design 3: 0.75x intraoral 2x vertical area wall, 15-degree curve. Design 4: No intraoral, 2x vertical area wall, 15-degree curve. Design 5: 0.75x intraoral, 0.5x vertical area wall, 20-degree curve. Design 6: 0.75x intraoral, 0.5x vertical area wall, 20-degree curve. Design 7: 0.75x intraoral, 0.5x vertical area wall, 15-degree curve. Design 8: 0.75x intraoral, 1.5x vertical area wall. 15-degree curve. Images of patient with iteration phase 5 designs. Images 1 and 2 show designs 1 and 5, respectively. Description: This iteration of prototypes aimed to further reduce the intrusiveness of the intra-oral piece, increase the surface area of the support wall, and further experiment with the angle of the extra-oral flap. The patient reported that the reduced volume of the intra-oral piece was more comfortable, and that a device without the intra-oral component would be ideal.

Phase 6: Varying intraoral piece, flat and convex wall support flaps

Iteration Phase 6 Designs: Design 1: No intraoral, 45-degree flat wall Design 2: small intraoral, 45-degree flat wall, small area Design 3: No intraoral, 50-degree flat wall Design 4: No intraoral, 30-degree flat wall Design 5: No intraoral, 30-degree convex wall Design 5: No intraoral, 50-degree convex wall Design 6: No intraoral, 45-degree convex wall. Images of patient with iteration phase 6 designs. Images 1 and 2 show designs 6 and 4, respectively. Description: This iteration of designs aimed to take out the intra-oral component, increase the rigidity of the support rod that held the wall support flap to increase the pressure of the device against the patient’s face, and to increase the area of the flap along the patients face. The patient found that this version helped him make a slightly better seal but was uncomfortable due to the large area of contact on his face.

Phase 7: Dynamic Spring- Assisted Hinged Unilateral Flap

Iteration Phase 7 Design. Front, top and isometric view of a dynamic device that uses springs to adjust the wall support angle to better conform to the patient’s changing embouchure. Images of patient with iteration phase 7 design from side (1) view and bottom (2) view. Description: The dynamic spring-assisted hinged unilateral flap prototype utilized a more dynamic approach, with ball and socket joints as well as a spring-loaded Y-arm to bend and move, conforming with the patients face. To increase comfort, the wall support portion of this device was made entirely of EcoFlex silicone rather than TPU. This device worked very well, but ultimately did not provide as much counter-pressure and support as the final device, which utilized a rigid wall construction.

Table S2. Final device-material cost breakdown.

| Material Name | Cost | Amount Used | Cost per device |
| --- | --- | --- | --- |
| Copper3D (Antimicrobial) MD FLEX TPU - Grey Blue - 2.85mm - 0.5KG | $67.05 / kg | 16.5 g | $1.11 per device |
| Ecoflex 00-30 - Super-Soft, Addition Cure Silicone Rubber | $40.32 / L | 20 g | $2.10 per device |
| Hose clamp (1) | $2.18 per clamp | 1 clamp | $2.18 |
| Zipties | $0.05 per zip tie | 2 zip ties | $0.10 |
| TOTAL COST |  |  | $5.49 |